PROJECT IMPLEMENTATION MANUAL

EMG-Controlled 3D-Printed Prosthetic Hand

with Myoelectric Actuators

Submitted by

Priyanshu Mohanty (23070127045)

Tej Datir (23070127062)

Ayush Chhaparia (23070127015)

Shounak Chaudhari (23070127070)

Guide: Dr. Satish Kumar

Co-Guide: Dr. Ramesh B.T

Department of Robotics and Automation

Symbiosis Institute of Technology, Pune

Group ID: 3  |  Batch 2023–27

2025–26

# 1. Project Overview

This manual provides step-by-step instructions to set up, assemble, wire, program, and evaluate the EMG-Controlled 3D-Printed Prosthetic Hand with Myoelectric Actuators. The system captures surface electromyography (sEMG) signals from residual limb muscles using a MyoWare 2.0 sensor, processes them in real time on an ESP32 microcontroller, and actuates servo motors to produce functional grasping motions via a tendon-driven finger mechanism.

| Parameter | Details |
| --- | --- |
| Project Title | EMG-Controlled 3D-Printed Prosthetic Hand with Myoelectric Actuators |
| Group ID | Group 3 |
| Microcontroller | ESP32 (Dual-core, 240 MHz, 12-bit ADC) |
| EMG Sensor | MyoWare 2.0 (sEMG, analog output) |
| Actuation | MG90S (fingers) + MG996R (shoulder, elbow, wrist) servo motors |
| Control Method | Threshold-Based EMG Control with PWM servo actuation |
| Structure | PLA+ via FDM 3D Printing (Creality K1 Max) |
| Power Supply | 3S LiPo (11.1V, 2200mAh) + LM2596 Buck Converters (6V & 5V rails) |
| Total Cost (Est.) | ₹13,000 – ₹15,750 INR (approx. USD 150–180) |
| Guide | Dr. Satish Kumar (Professor, PhD In-charge) |
| Co-Guide | Dr. Ramesh B.T (Assistant Professor) |

# 2. System Architecture

The prosthetic system is structured in four functional layers, each responsible for a distinct aspect of operation:

- Sensing Layer – MyoWare 2.0 EMG sensor placed on the forearm flexor muscle captures sEMG signals and outputs an analog voltage proportional to muscle activation amplitude.

- Processing Layer – ESP32 reads the analog signal via its 12-bit ADC (GPIO 34), applies a 10-sample rolling moving-average filter to reduce noise, and compares the processed value against a user-calibrated threshold with ±10% hysteresis deadband.

- Actuation Layer – When the threshold is crossed, the ESP32 generates PWM signals (50 Hz, 1–2 ms pulse width) to drive the MG90S finger servos and MG996R joint servos to preset grip positions.

- Power Layer – The 3S LiPo battery feeds two LM2596 DC-DC converters: one trimmed to 6V for servo motors and one trimmed to 5V for ESP32 and sensor logic. Total estimated runtime is 2–4 hours under normal duty cycle.

# 3. Bill of Materials

| # | Component | Model / Specification | Qty | Approx. Cost (INR) |
| --- | --- | --- | --- | --- |
| 1 | Microcontroller | ESP32 Dev Board (38-pin) | 1 | ₹380 |
| 2 | EMG Sensor | MyoWare 2.0 (Analog sEMG) | 1 | ₹4,000 |
| 3 | Cable Shield (EMG) | MyoWare cable shield + patches | 1 set | ₹800 |
| 4 | Finger Servo Motors | MG90S Micro Servo (2.5 kg·cm, 14.7 g) | 5 | ₹1,323 |
| 5 | Joint Servo Motors | MG996R High-Torque (9.4 kg·cm @ 4.8V) | 3 | ₹1,408 |
| 6 | Power Battery | 3S LiPo 11.1V 2200mAh | 1 | ₹1,500 |
| 7 | Buck Converters | LM2596 DC-DC Step-Down Module | 2 | ₹299 |
| 8 | Logic Level Converter | 3.3V – 5V Bi-directional | 1 | ₹20 |
| 9 | Connectors | XT60 Battery Connectors + Wiring Harness | 1 set | ₹200 |
| 10 | PLA+ Filament | Generic PLA+ (220°C nozzle, 60°C bed) | ~800 g | ₹2,000 |
| 11 | Braided Fishing Line | Flexor tendons (monofilament, routed through conduit channels) | 5 m | ₹150 |
| 12 | Elastic Orthodontic Bands | Extensor return mechanism (dorsal side of each finger) | 10 pcs | ₹50 |
| 13 | M3/M4 Hardware Fasteners | Screws, heat-staked inserts for assembly | 1 set | ₹200 |
| 14 | Backup Servo Motors | Extra MG90S/MG996R for spares | 2 | ₹2,500 |
| 15 | Miscellaneous (wires, tape, etc.) | Jumper wires, heat shrink, insulation tape | 1 set | ₹220 |

Estimated Total: ₹15,750 INR (approximately USD 180)

# 4. Mechanical Design & 3D Printing

All structural components were designed in Fusion 360 and sliced using OrcaSlicer for printing on the Creality K1 Max FDM printer. The prosthetic spans a complete upper-limb skeleton from shoulder to fingertip, divided into independently replaceable modules.

## 4.1 Modular Assembly Structure

- Shoulder Module – MG996R servo achieves flexion / extension (shoulder frame + upper arm casing).

- Elbow Module – MG996R servo achieves flexion / extension (elbow joint part + lower arm housing).

- Wrist Module – MG996R servo achieves pronation / supination (wrist connector + palm optimized connector).

- Hand & Finger Module – Five MG90S micro-servos housed in the palm plate, each attached to a tendon-driven finger assembly. Braided fishing line routes through printed conduit channels acting as flexor tendons; elastic orthodontic bands on the dorsal surface act as passive extenders.

## 4.2 3D Printing Settings

| Print Parameter | Value |
| --- | --- |
| Filament | Generic PLA+ |
| Nozzle Temperature | 220°C |
| Bed Temperature | 60°C |
| Layer Height | 0.2 mm |
| Infill Density | 40% Gyroid |
| Print Plate Type | Smooth High Temperature Print Plate |
| Total Print Plates | 13 plates |
| Total Print Time | ~38 hours |
| Dimensional Tolerance | ±0.3 mm on critical mating surfaces |

## 4.3 Post-Processing & Assembly

- Sand and deburr all critical seat areas using fine-grit sandpaper after printing.

- Remove any bridged material from conduit channels using a thin rod or wire.

- Insert M3/M4 screws or heat-staked threaded inserts at all joint connection points.

- Route braided fishing line from servo spool mounting points through the finger conduit channels to each fingertip.

- Affix elastic orthodontic bands to the dorsal surface of each finger and anchor to the palmar side of the hand chassis.

- Verify full open/close range of motion manually (pulling tendons and releasing) before installing servos.

# 5. Electronics Wiring & Integration

## 5.1 EMG Sensor to ESP32

The MyoWare 2.0 sensor shielded output wire connects directly to ESP32 GPIO 34 (ADC1_CH6, 12-bit). The sensor operates at 5V; no logic-level conversion is needed on the analog output since the ESP32 ADC input is tolerant up to 3.3V — ensure the MyoWare SIG output does not exceed 3.3V at rest, or insert a voltage divider (10kΩ / 10kΩ) inline.

| MyoWare 2.0 Pin | Connection | ESP32 / Power |
| --- | --- | --- |
| VCC (+) | — | 5V logic rail |
| GND | — | GND (common ground) |
| SIG (shielded output) | Direct / voltage divider | GPIO 34 (ADC1_CH6) |

## 5.2 PCA9685 PWM Driver to ESP32 (I2C)

All servo motors are driven through a PCA9685 16-channel PWM driver board connected to the ESP32 via I2C. This frees the ESP32 from generating multiple simultaneous PWM signals and allows up to 16 channels at precise 12-bit resolution.

| PCA9685 Pin | ESP32 GPIO | Power | Notes |
| --- | --- | --- | --- |
| SDA | GPIO 21 | — | I2C Data |
| SCL | GPIO 22 | — | I2C Clock |
| VCC | — | 5V logic rail | Board logic power |
| V+ (servo power) | — | 6V servo rail | Servo motor power (all channels) |
| GND | — | Common GND | Shared ground with ESP32 |

## 5.3 Servo Channel Mapping (PCA9685)

Finger groups are consolidated onto three PCA9685 channels. Note that channel 1 (middle finger) uses mechanically reversed servo mounting — the firmware corrects this by inverting the pulse value automatically (finalVal = F_MIN + F_MAX − val).

| PCA9685 Ch | Finger Group | Motor | Open (ticks) | Closed (ticks) |
| --- | --- | --- | --- | --- |
| 0 | Thumb + Index | MG90S | 150 (F_MIN) | 600 (THUMB_MAX) |
| 1 | Middle (reversed) | MG90S | 500 (F_MAX) | 150 (F_MIN) |
| 2 | Ring + Pinky | MG90S | 150 (F_MIN) | 500 (F_MAX) |
| 3 | Wrist | MG996R | 307 (STOP_VAL) | 307 (neutral) |
| 4 | Elbow | MG996R | 307 (STOP_VAL) | 307 (neutral) |

## 5.4 Power Circuit

- 3S LiPo (11.1V) → LM2596 Buck Converter A → 6.00V servo rail (for all MG90S + MG996R motors).

- 3S LiPo (11.1V) → LM2596 Buck Converter B → 5.00V logic rail (for ESP32, MyoWare sensor, logic-level converter).

- Trim output voltages under no-load conditions first, then verify under simultaneous load of five MG90S servos (typical draw: 300–600 mA per servo).

- Connect battery via XT60 connector with an inline power switch for safe operation.

- Route all cables through the dedicated cable channels built into the forearm housing to avoid tangling with the tendon mechanism.

# 6. Software & Firmware Setup

## 6.1 Development Environment

IDE: Arduino IDE 2.x with ESP32 board support package. The ESP32 hosts a Wi-Fi access point (SSID: "Bionic-Arm-Pro") — connect any device to this network and open http://192.168.4.1 to access the live dashboard.

| Library | Purpose |
| --- | --- |
| WiFi.h | ESP32 Wi-Fi softAP — creates the Bionic-Arm-Pro hotspot |
| WebServer.h | HTTP server on port 80 — dashboard page + REST endpoints (/set, /getEMG, /toggleEMG) |
| Wire.h | I2C communication with PCA9685 PWM driver (SDA=GPIO21, SCL=GPIO22) |
| Adafruit_PWMServoDriver.h | Controls all servo channels via PCA9685; 12-bit resolution at 50 Hz |

## 6.2 Complete Firmware

Flash the following sketch to the ESP32. All tunable parameters (threshold, pulse ranges, Wi-Fi credentials) are defined as constants at the top for easy adjustment without modifying the main logic.

#include <WiFi.h>

#include <WebServer.h>

#include <Wire.h>

#include <Adafruit_PWMServoDriver.h>

const char* ssid = "Bionic-Arm-Pro";

const char* password = "password123";

Adafruit_PWMServoDriver pwm = Adafruit_PWMServoDriver();

WebServer server(80);

// Pins & Calibration

const int EMG_PIN = 34;

#define SERVO_FREQ 50

#define STOP_VAL 307

// Finger Pulse Ranges

#define F_MIN 150

#define F_MAX 500

#define THUMB_MAX 600

// EMG Variables

float filteredValue = 0;

bool emgActive = false;

int userThreshold = 800; // Unified threshold for all finger motors

void handleRoot() {

String html = "<!DOCTYPE html><html><head><meta name='viewport' content='width=device-width, initial-scale=1.0'>";

html += "<script src='https://cdn.jsdelivr.net/npm/chart.js'></script>";

html += "<style>body{font-family:sans-serif; text-align:center; background:#121212; color:#00ffcc; padding:10px;}";

html += ".slider{width:90%; height:25px; margin:10px 0;} .btn{padding:15px; margin:10px; background:#ff3300; color:#fff; border:none; border-radius:8px; font-weight:bold; width:90%;}</style></head>";

html += "<body><h1>BIONIC HYBRID v2.1</h1>";

html += "<div style='width:100%; height:200px;'><canvas id='emgChart'></canvas></div>";

html += "<button id='toggleBtn' class='btn' onclick='toggleEMG()'>ENABLE EMG CONTROL</button>";

// Sliders for manual override

html += "<h3>Thumb/Index (0-290)</h3><input type='range' min='150' max='600' class='slider' oninput='update(0,this.value)'>";

html += "<h3>Middle (Reversed)</h3><input type='range' min='150' max='500' class='slider' oninput='update(1,this.value)'>";

html += "<h3>Ring/Pinky</h3><input type='range' min='150' max='500' class='slider' oninput='update(2,this.value)'>";

html += "<hr><h3>Wrist</h3><input type='range' min='200' max='400' value='307' class='slider' id='s3' oninput='update(3,this.value)' onmouseup='reset(3)'>";

html += "<h3>Elbow</h3><input type='range' min='200' max='400' value='307' class='slider' id='s4' oninput='update(4,this.value)' onmouseup='reset(4)'>";

html += "<script>var chartCtx = document.getElementById('emgChart').getContext('2d');";

html += "var chart = new Chart(chartCtx, {type:'line', data:{labels:[], datasets:[{label:'EMG', data:[], borderColor:'#00ffcc', fill:false, borderWidth:1}]}, options:{animation:false, scales:{y:{min:0, max:4095}}}});";

html += "function update(ch,v){ fetch('/set?ch='+ch+'&val='+v); }";

html += "function reset(ch){ fetch('/set?ch='+ch+'&val=307'); document.getElementById('s'+ch).value=307; }";

html += "function toggleEMG(){ fetch('/toggleEMG').then(r=>r.text()).then(t=>{ document.getElementById('toggleBtn').innerText = t; document.getElementById('toggleBtn').style.background = (t.includes('ENABLE'))?'#ff3300':'#00cc66'; }); }";

html += "setInterval(()=>{ fetch('/getEMG').then(r=>r.text()).then(v=>{ if(chart.data.labels.length>20){chart.data.labels.shift(); chart.data.datasets[0].data.shift();} chart.data.labels.push(''); chart.data.datasets[0].data.push(v); chart.update(); }); }, 100);</script></body></html>";

server.send(200, "text/html", html);

}

void handleSet() {

if (server.hasArg("ch") && server.hasArg("val")) {

int ch = server.arg("ch").toInt();

int val = server.arg("val").toInt();

int finalVal = val;

if (ch == 1) finalVal = F_MIN + F_MAX - val; // Reverse Middle logic

pwm.setPWM(ch, 0, finalVal);

server.send(200, "text/plain", "OK");

}

}

void setup() {

Serial.begin(115200);

Wire.begin(21, 22);

analogReadResolution(12);

pwm.begin();

pwm.setPWMFreq(SERVO_FREQ);

pwm.setPWM(3, 0, STOP_VAL);

pwm.setPWM(4, 0, STOP_VAL);

WiFi.softAP(ssid, password);

server.on("/", handleRoot);

server.on("/set", handleSet);

server.on("/getEMG", [](){ server.send(200, "text/plain", String((int)filteredValue)); });

server.on("/toggleEMG", [](){

emgActive = !emgActive;

server.send(200, "text/plain", emgActive ? "DISABLE EMG CONTROL" : "ENABLE EMG CONTROL");

});

server.begin();

}

void loop() {

server.handleClient();

// 1. Read and Smooth EMG

int rawValue = analogRead(EMG_PIN);

filteredValue = (filteredValue * 0.9) + (rawValue * 0.1);

// 2. SERIAL PLOTTER OUTPUT

// This allows you to see the threshold vs the actual signal

Serial.print("EMG_Signal:");

Serial.print(filteredValue);

Serial.print(",");

Serial.print("Threshold:");

Serial.println(userThreshold);

// 3. Unified Finger Control

if (emgActive) {

if (filteredValue > userThreshold) {

// CLOSE FIST

pwm.setPWM(0, 0, THUMB_MAX); // Thumb/Index to 290 deg pos

pwm.setPWM(1, 0, F_MIN);     // Middle (Reversed logic: F_MIN is closed)

pwm.setPWM(2, 0, F_MAX);     // Ring/Pinky to 180 deg pos

} else {

// OPEN FIST

pwm.setPWM(0, 0, F_MIN);     // Thumb/Index open

pwm.setPWM(1, 0, F_MAX);     // Middle (Reversed logic: F_MAX is open)

pwm.setPWM(2, 0, F_MIN);     // Ring/Pinky open

}

}

}

# 7. EMG Calibration Procedure

Calibration must be performed each time the sensor is repositioned. Follow these steps in order:

- Place the MyoWare 2.0 sensor on the wrist flexor digitorum superficialis muscle of the forearm (distal aspect of the radius). Ensure the three electrode pads are aligned along the muscle belly.

- Power the system on. Open the Arduino IDE Serial Plotter at 115,200 baud. You will see two live traces: EMG_Signal (EMA-filtered ADC value) and Threshold (the current userThreshold, default 800).

- Relax the arm completely. Observe the resting EMG_Signal trace and note the settled mean ADC value at rest (typically 100–200 ADC counts with the EMA filter settled).

- Perform a firm wrist flexor contraction (squeeze the fist). Observe the peak EMG_Signal value during contraction (typically 800–3200 ADC counts depending on electrode placement and individual muscle strength).

- Set userThreshold in the firmware to a value midway between the mean rest and mean contraction values. Re-flash the ESP32 if needed.

- The EMA filter (coefficient 0.9 retain / 0.1 new sample) provides inherent noise rejection — brief noise spikes are attenuated naturally without requiring a separate hysteresis variable.

- Confirm calibration: perform 20 supervised grip trials targeting ≥18 correct triggers with ≤2 false positives. If false positives persist, raise userThreshold by 50 ADC counts and retest.

# 8. Deployment Steps

Follow these steps in sequence for a clean first system deployment:

- Flash the ESP32 firmware via Arduino IDE. Open the Serial Plotter at 115,200 baud and verify EMG_Signal and Threshold traces appear correctly before proceeding.

- Trim both LM2596 buck converters under no-load: Converter A to 6.00V (servo rail), Converter B to 5.00V (logic rail). Re-verify voltages under simultaneous PCA9685 + five-servo load.

- Connect PCA9685 board to ESP32 via I2C (SDA=GPIO21, SCL=GPIO22). Connect all servo signal wires to the correct PCA9685 channels as per Section 5.3.

- Mount the ESP32, PCA9685, and both LM2596 buck converters inside the forearm housing on the printed mounting bracket. Route all cables through the built-in cable channels.

- Connect the 3S LiPo battery via the XT60 connector. Power on the system and verify wrist (ch3) and elbow (ch4) initialize to neutral position (STOP_VAL = 307 ticks).

- From a phone or PC, connect to Wi-Fi network "Bionic-Arm-Pro" (password: password123) and open http://192.168.4.1. Verify the dashboard loads and the live EMG chart is updating.

- Attach MyoWare 2.0 sensor with cable shield to the forearm. Perform EMG calibration as per Section 7. Verify threshold in Serial Plotter output.

- On the dashboard, press ENABLE EMG CONTROL. Perform 20 supervised grip trials and confirm ≥90% correct trigger rate before final deployment.

- Mount the assembled prosthetic on the physical test rig and secure the shoulder socket frame. Run full-arm range-of-motion check via dashboard manual sliders.

# 9. Configuration Reference

| Parameter | Default Value | Location to Change |
| --- | --- | --- |
| EMG ADC Pin | GPIO 34 | Firmware – EMG_PIN define |
| EMG Threshold (userThreshold) | 800 (ADC counts, 0–4095) | Firmware – userThreshold variable; adjust via Serial Plotter calibration |
| EMA Filter Coefficient | 0.9 (retain) / 0.1 (new) | Firmware – filteredValue formula in loop() |
| Wi-Fi SSID | Bionic-Arm-Pro | Firmware – ssid variable |
| Wi-Fi Password | password123 | Firmware – password variable |
| Dashboard URL | http://192.168.4.1 | Fixed (ESP32 softAP default gateway) |
| PWM Frequency | 50 Hz | Firmware – SERVO_FREQ define |
| Wrist / Elbow Neutral (STOP_VAL) | 307 ticks | Firmware – STOP_VAL define |
| Finger Open (F_MIN) | 150 ticks | Firmware – F_MIN define |
| Finger Closed (F_MAX) | 500 ticks | Firmware – F_MAX define |
| Thumb/Index Closed (THUMB_MAX) | 600 ticks | Firmware – THUMB_MAX define |
| I2C Pins (PCA9685) | SDA=GPIO21, SCL=GPIO22 | Firmware – Wire.begin(21, 22) |
| Serial Baud Rate | 115,200 baud | Firmware – Serial.begin(115200) |
| Servo Rail Voltage | 6.00V | LM2596 Converter A trim potentiometer |
| Logic Rail Voltage | 5.00V | LM2596 Converter B trim potentiometer |

# 10. Performance Evaluation & Testing

The following test protocol was used to evaluate system performance on the physical test rig. Conduct a minimum of 30 trials per grip type (power grasp and lateral pinch grasp) over 5 evaluation sessions.

| Evaluation Metric | Measurement Method | Target / Observed Result |
| --- | --- | --- |
| Grip Response Time (L_total) | Oscilloscope / High-speed camera from EMG threshold crossing to full grip closure | 55–160 ms (est.); observed 120–180 ms |
| Control Accuracy | % of voluntary contractions that produced the correct grip pattern vs. false triggers | Target ≥90%; achieved 90% in 20 trials |
| EMG Signal Reliability (SNR) | ADC count ratio: rest vs. contraction via Serial Plotter | 15–20 dB S/N ratio observed |
| Finger Extension Return Time | Time for elastic bands to return fingers to neutral after grip release | <200 ms consistently observed |
| Wrist/Elbow Range of Motion | Angle sweep time via PWM ramp test | Full ±90° range in 200–250 ms |
| Structural Integrity | Visual inspection under simulated gripping forces | No deformation or delamination observed |
| System Repeatability | Consistency across 100+ grip trials over multiple sessions | No slippage or creep in braided tendon anchors |

# 11. Troubleshooting

| Issue | Likely Cause | Fix |
| --- | --- | --- |
| Servo does not move on contraction | EMG threshold too high or EMG signal too low | Recalibrate threshold; reposition MyoWare sensor closer to muscle belly |
| False triggers at rest | Threshold too low or motion artifact noise | Increase threshold; increase hysteresis band to ±15%; check cable shield connection |
| EMG ADC reading is 0 or constant | Sensor power or ADC pin issue | Verify 5V rail voltage; check GPIO 34 connection and logic-level converter |
| Finger does not extend back after grip | Elastic band broken or insufficient tension | Replace elastic orthodontic band; check dorsal anchor attachment point |
| Servo jitter / oscillation | PWM noise or insufficient supply voltage | Verify 6V servo rail under load; add 100µF capacitor across servo power pins |
| ESP32 resets under load | Brownout from LiPo sag under five-servo load | Verify LM2596 voltage stability; check LiPo charge level; add 470µF bulk cap |
| Tendon slips or finger does not fully close | Braided line anchor slippage on servo spool | Re-knot braided line anchor; apply small drop of CA glue to spool wrap |
| Servo overheats during operation | Continuous stall current or wrong voltage | Check that 6V rail does not exceed 6.2V; reduce grip hold duration; verify servo range of motion does not stall |
| Wrist / elbow joint binds mechanically | Misalignment or insufficient joint clearance after printing | Sand mating surfaces lightly; verify M3 insert alignment; check torque arm bracket orientation |

# 12. System Latency Model

The total system latency (L_total) is the sum of three components:

| Component | Description | Typical Value |
| --- | --- | --- |
| L_emg | Signal conditioning delay in MyoWare 2.0 sensor (RMS envelope extraction) | <5 ms |
| L_processing | ESP32 ADC sampling + rolling average filter computation at 240 MHz | <2 ms at 500 Hz sampling |
| L_servo | Mechanical response time of MG90S servo to achieve target position (0.1 s per 60° at 5V) | 120–180 ms (full close) |
| L_total | L_emg + L_processing + L_servo | Estimated 127–187 ms |

The dominant latency component is the mechanical servo response. L_total is within acceptable human perception thresholds for assistive prosthetic control.

# 13. Future Enhancements

- Integrate machine learning gesture classification (SVM or shallow neural network) on the ESP32 to enable multiple grip patterns without hardware changes.

- Add FSR (Force Sensitive Resistor) fingertip sensors to implement closed-loop grip force feedback control.

- Implement vibrotactile haptic feedback using mini motors attached to the residual limb for grip force information.

- Explore carbon fiber reinforced PETG or TPU flexible materials for the socket portion to improve durability, comfort, and aesthetic quality.

- Test system functionality with volunteer amputees through collaboration with a prosthetic and rehabilitation clinic.

- Implement multi-modal sensing by fusing EMG with IMU (inertial measurement unit) data for improved gesture recognition accuracy and context awareness.

- Develop OTA (Over-the-Air) firmware update capability via ESP32 Wi-Fi for remote calibration and parameter updates.

Symbiosis Institute of Technology, Pune  |  Department of Robotics and Automation  |  Batch 2023–27